T, S, DO, PH

- Hydrolab arrived following factory repairs and was used for sampling

  - Had some issues leading to instrument failure and in process for factory exchange

  - Some S, pH, and DO produced illogically high/low values that were removed from the data and appear as NA

- Backup YSI required factory calibration and did not produce useable data

- Eureka Sonde arrived but had issue with salinity scaling that required factory service, was serviced during Quarter 2, and was used for remainder of sites and will be primary sonde for future surveys

- Samples were preserved for lab analysis, but were discarded before being analyzed by a calibrated conductivity meter

  - Courtney discussed this issue with chemistry lab staff to avoid future issues

TSS

- TSS 462 was held for 50 hrs 1 mins prior to analysis

- TSS 352 was held for 48 hours 28 mins prior to analysis

- Standard within 5 mg and blanks <= 0.60 mg/L

- Everything else looks good

Enterococcus

- Everything looks good

- Batch quality control checks previously completed in spring 2023

Suspended Solids

- SS 461 exceeded holding time by 2 hrs and 31 mins

- SS 351 exceeded holding time by 53 mins

- Everything else looks good

BOD

- 2.0+-0.8 mg/L GGA spike ranged from 1.31 to 1.98 mg/L within QAPP

- BOD 260 0.39 and 0.83 mg/L outside of +-0.2 mg/L range by 0.04 mg/L

- All other BOD duplicates within +- 0.2 mg/L

Chlorophyll-a

- Standards meet QA/QC checks

Turbidity

- 460 exceeded holding time by 8 hours

- 350 exceeded holding time by 6 hours 24 mins

- 449 exceeded holding time by 5 hours 28 mins

- All else appears correct

DIC

- 9 surface and 5 bottom samples were lost due to a malfunction of the DIC analyzer

- DIC left in units of uM owing to missing salinity values required to calculate umol/kg

TA

- Samples were incorrectly analyzed with respect to CRM and were discarded before they could be re-analyzed

  - Courtney discussed this issue with chemistry lab staff to avoid future issues

NOx

- Some samples exceeded handling time while troubleshooting errors with the AA500 Nutrient AutoAnalyzer

- 10/10/2023

  - Correlation coefficient of standards = 0.9997

  - CRM QC standard passed

  - Spike recovery was 101%, 103% and 18%

- 11/9/2023

  - Correlation coefficient of standards > 0.9999

  - CRM QC standard passed

  - Spike recovery was 92%, 96%, 84%, and 81%

o-PO4

- Some samples exceeded handling time while troubleshooting errors with the AA500 Nutrient AutoAnalyzer

- 10/10/2023

  - Correlation coefficient of standards = 0.9999

  - Incorrect CRM for o-PO4 so none to report

    - New CRM ordered for Q3

  - Spike recovery was 99%, 104% and 99%

- 11/9/2023

  - Correlation coefficient of standards > 0.9999

  - Incorrect CRM for o-PO4 so none to report

    - New CRM ordered for Q3

  - Spike recovery was 92%, 94%, 92%, 84%, and 90%

TKN

- Some samples exceeded handling time while troubleshooting errors with the AA500 Nutrient AutoAnalyzer

- 11/22/2023

  - Correlation coefficient of standards = 0.9997

  - CRM (6.0 mg/) ranged from 5.686 to 6.123 (PASS)

  - Spike recovery was 103%, 106%, 116%, 101%, 28%, -34%, 476%, -37%, 29%

    - Unclear why some spikes pass (100%) and why some were quite off 😣