| Date | Time | User |  | Sonde |  |
| --- | --- | --- | --- | --- | --- |
| February 2nd, 2024 | 4:41 am | Manuel A. Nieves Ortiz |  | Eureka Manta |  |
| Parameter | Test | Standard | Measured | Pass/Fail | Note |
| Depth | Check | 0.00 | -0.10 | P |  |
| (± 0.4 m) | Calibration |  |  |  |  |
|  | Calib. Check |  |  |  |  |
| Temperature | Check | 24.20 | 24.16 | P |  |
| (± 0.1°C) | Calibration |  |  |  |  |
|  | Calib. Check |  |  |  |  |
| Sp. Conductivity | Check | 50.00 | 49.89 | P |  |
| (± 0.5% or ± 0.25 mS/cm for a 50 mS/cm std) | Calibration |  |  |  |  |
|  | Calib. Check |  |  |  |  |
| pH | Check | 9.00 | 9.06 | P |  |
| (± 0.1 units) | Calibration |  |  |  |  |
|  | Calib. Check |  |  |  |  |
| Dissolved Oxygen | Check | 100.0 | 98.4 | F |  |
| (± 0.1 mg/L or ± 1.5 %) | Calibration | 100.0 | 98.2 | F |  |
|  | Calib. Check | 100.0 | 99.8 | P |  |

Notes: