| Date | Time | User |  | Sonde |  |
| --- | --- | --- | --- | --- | --- |
| February 15th, 2023 | 5:36 am | Manuel A. Nieves Ortiz |  | Eureka Manta |  |
| Parameter | Test | Standard | Measured | Pass/Fail | Note |
| Depth | Check | 0.00 | -0.02 | P |  |
| (± 0.4 m) | Calibration |  |  |  |  |
|  | Calib. Check |  |  |  |  |
| Temperature | Check | 23.07 | 23.06 | P |  |
| (± 0.1°C) | Calibration |  |  |  |  |
|  | Calib. Check |  |  |  |  |
| Sp. Conductivity | Check | 50.00 | 50.03 | P |  |
| (± 0.5% or ± 0.25 mS/cm for a 50 mS/cm std) | Calibration |  |  |  |  |
|  | Calib. Check |  |  |  |  |
| pH | Check | 9.00 | 9.01 | P |  |
| (± 0.1 units) | Calibration |  |  |  |  |
|  | Calib. Check |  |  |  |  |
| Dissolved Oxygen | Check | 100.0 | 102.1 | F |  |
| (± 0.1 mg/L or ± 1.5 %) | Calibration | 100.0 | 101.7 | F |  |
|  | Calib. Check | 100.0 | 99.7 | P |  |

Notes: