| Date | Time | User |  | Sonde |  |
| --- | --- | --- | --- | --- | --- |
| February 21st 2024 | 5:49 am | Manuel A. Nieves Ortiz |  | Eureka Manta |  |
| Parameter | Test | Standard | Measured | Pass/Fail | Note |
| Depth | Check | 0.00 | -0.06 | P |  |
| (± 0.4 m) | Calibration |  |  |  |  |
|  | Calib. Check |  |  |  |  |
| Temperature | Check | 21.51 | 21.54 | p |  |
| (± 0.1°C) | Calibration |  |  |  |  |
|  | Calib. Check |  |  |  |  |
| Sp. Conductivity | Check | 50.00 | 50.07 | P |  |
| (± 0.5% or ± 0.25 mS/cm for a 50 mS/cm std) | Calibration |  |  |  |  |
|  | Calib. Check |  |  |  |  |
| pH | Check | 9.00 | 9.04 | P |  |
| (± 0.1 units) | Calibration |  |  |  |  |
|  | Calib. Check |  |  |  |  |
| Dissolved Oxygen | Check | 100.0 | 98.1 | F |  |
| (± 0.1 mg/L or ± 1.5 %) | Calibration | 100.0 | 98.8 | F |  |
|  | Calib. Check | 100.0 | 100.1 | P |  |

Notes: