| Date | Time | User |  |  |  | Sonde |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Jan 22, 2024 | 7:11 | Manuel Nieves |  |  |  | Eureka Manta |  |  |
| Parameter | Test | Standard |  | Measured |  | Pass/Fail |  | Note |
| Depth | Check | 0.0 |  | 0.06 |  | Pass |  |  |
| (± 0.4 m) | Calibration |  |  |  |  |  |  |  |
|  | Calib. Check | 0.0 |  | 0.0 |  | Pass |  |  |
| Temperature | Check | 24.14 |  | 24.00 |  | Fail |  |  |
| (± 0.1°C) | Calibration |  |  |  |  |  |  |  |
|  | Calib. Check | 24.18 |  | 24.17 |  | Pass |  | Hold thermometer over sensors |
| Sp. Conductivity | Check | 50 |  | 49.86 |  | Pass |  |  |
| (± 0.5% or ± 0.25 mS/cm for a 50 mS/cm std) | Calibration |  |  |  |  |  |  |  |
|  | Calib. Check |  |  |  |  |  |  |  |
| pH | Check | 9.0 |  | 8.84 |  | Fail |  |  |
| (± 0.1 units) | Calibration | 9.0 | 8.82 | 7.0 | 7.02 | F | P |  |
|  | Calib. Check | 7.0 |  | 6.99 |  | Pass |  |  |
| Dissolved Oxygen | Check | 100 |  | 99.4 |  | Pass |  |  |
| (± 0.1 mg/L or ± 1.5 %) | Calibration |  |  |  |  |  |  |  |
|  | Calib. Check |  |  |  |  |  |  |  |

Notes: Hold thermometer over sensors when doing temperature check