| Date | Time | User |  | Sonde |  |
| --- | --- | --- | --- | --- | --- |
| January 31st, 2023 | 4:33 am | Manuel A. Nieves Ortiz |  | Eureka Manta |  |
| Parameter | Test | Standard | Measured | Pass/Fail | Note |
| Depth | Check | 0.00 | 0.08 | P |  |
| (± 0.4 m) | Calibration |  |  |  |  |
|  | Calib. Check |  |  |  |  |
| Temperature | Check | 26.51 | 26.51 | P |  |
| (± 0.1°C) | Calibration |  |  |  |  |
|  | Calib. Check |  |  |  |  |
| Sp. Conductivity | Check | 50.00 | 49.90 | P |  |
| (± 0.5% or ± 0.25 mS/cm for a 50 mS/cm std) | Calibration |  |  |  |  |
|  | Calib. Check |  |  |  |  |
| pH | Check | 9.00 | 9.03 | P |  |
| (± 0.1 units) | Calibration |  |  |  |  |
|  | Calib. Check |  |  |  |  |
| Dissolved Oxygen | Check | 100.0 | 98.9 | P |  |
| (± 0.1 mg/L or ± 1.5 %) | Calibration |  |  |  |  |
|  | Calib. Check |  |  |  |  |

Notes: