| Date | Time | User |  |  |  | Sonde |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| May 3rd, 2024 | 6:50 | Manuel A. Nieves Ortiz |  |  |  | Eureka |  |
| Parameter | Test | Standard |  | Measured |  | Pass/Fail | Note |
| Depth | Check | 0.00 |  | -0.04 |  | P |  |
| (± 0.4 m) | Calibration |  |  |  |  |  |  |
|  | Calib. Check |  |  |  |  |  |  |
| Temperature | Check | 26.73 |  | 26.73 |  | P |  |
| (± 0.1°C) | Calibration |  |  |  |  |  |  |
|  | Calib. Check |  |  |  |  |  |  |
| Sp. Conductivity | Check | 50.00 |  | 50.27 |  | F |  |
| (± 0.5% or ± 0.25 mS/cm for a 50 mS/cm std) | Calibration | 50.00 |  | 50.32 |  | F |  |
|  | Calib. Check | 50.00 |  | 49.98 |  | P |  |
| pH | Check | 10.00 |  | 10.20 |  | F |  |
| (± 0.1 units) | Calibration | 10.00 | 7.00 | 10.20 | 6.91 | F |  |
|  | Calib. Check | 10.00 |  | 9.99 |  | P |  |
| Dissolved Oxygen | Check | 100.0 |  | 99.1 |  | P |  |
| (± 0.1 mg/L or ± 1.5 %) | Calibration |  |  |  |  |  |  |
|  | Calib. Check |  |  |  |  |  |  |

Notes: