| Date | Time | User |  |  |  | Sonde |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| May 5th, 2024 | 4:45 am | Manuel A. Nieves Ortiz |  |  |  | Eureka |  |  |
| Parameter | Test | Standard |  | Measured |  | Pass/Fail |  | Note |
| Depth | Check | 0.00 |  | -0.05 |  | P |  |  |
| (± 0.4 m) | Calibration |  |  |  |  |  |  |  |
|  | Calib. Check |  |  |  |  |  |  |  |
| Temperature | Check | 26.74 |  | 26.74 |  | P |  |  |
| (± 0.1°C) | Calibration |  |  |  |  |  |  |  |
|  | Calib. Check |  |  |  |  |  |  |  |
| Sp. Conductivity | Check | 50.00 |  | 49.89 |  | P |  |  |
| (± 0.5% or ± 0.25 mS/cm for a 50 mS/cm std) | Calibration |  |  |  |  |  |  |  |
|  | Calib. Check |  |  |  |  |  |  |  |
| pH | Check | 10.00 |  | 10.23 |  | F |  |  |
| (± 0.1 units) | Calibration | 10.00 | 7.00 | 10.23 | 6.95 | F | P |  |
|  | Calib. Check | 7.00 |  | 7.00 |  | P |  |  |
| Dissolved Oxygen | Check | 100.0 |  | 99.4 |  | P |  |  |
| (± 0.1 mg/L or ± 1.5 %) | Calibration |  |  |  |  |  |  |  |
|  | Calib. Check |  |  |  |  |  |  |  |

Notes: